EMC Test Plan issued under the responsibility of:

Company Name:

Address:

| EMC TEST PLAN IEC 60601-1-2 Medical Electrical Equipment PART 1-2: General Requirements for Basic Safety and Essential Performance Collateral Standard: Electromagnetic Compatibility |
| --- |
| Test Plant Number. : |  |
| Date of issue : |  |
| Total number of pages |  |
| Name of Person preparing the Test Plan : |  |
| Applicant’s name : |  |
| Address : |  |
| Test specification: |  |
| Standard : | IEC 60601-1-2: 2014 (Fourth Edition) |
| Test procedure : | IEC 60601-1-2: 2014 (Fourth Edition) |
| Non-standard test method : | N/A |
| General disclaimer: |
| The test plan presented in this report relate only to the object to be tested. This test plan shall not be reproduced, except in full, without the written approval of the Issuing agent. |

| Test item description : |  |
| --- | --- |
| Trade Mark : |  |
| Manufacturer : |  |
| Model/Type reference : |  |
| Ratings : |  |
| Responsible Testing Laboratory (as applicable), testing procedure and testing location(s): |
|  | Testing Laboratory: |  |
| Testing location/ address : |  |
|  | Associated Testing Laboratory: |  |
| Testing location/ address : |  |

| Test item particulars : |  |
| --- | --- |
| Classification of installation and use : |  |
| Supply Connection : |  |
| : |  |
| Possible test case verdicts: |  |
| - test case does not apply to the test object : | N/A |
| - test object does meet the requirement : | P (Pass) |
| - test object does not meet the requirement : | F (Fail) |
| Testing : |  |
| Date of receipt of test item : |  |
| Date (s) of performance of tests : |  |
| General remarks: |
| "(See Enclosure #)" refers to additional information appended to the report. "(See appended table)" refers to a table appended to the report. Throughout this report a comma / point is used as the decimal separator. |
| When differences exist; they shall be identified in the General product information section. |
| Name and address of factory (ies) : |  |
| General product information: |
| Software and Firmware Version: |
| Prototype : Production Version: |
| Unit(s) Tested (include serial numbers): Rationale for number of samples tested: |
| Intended Use: |
| Intended healthcare environments (Specify environment: Professional, Home, Aircraft, Special or Other (if other please describe)): |
| Any deviations from the Basic EMC standards or from this collateral standard: |

| Test Plan Index: |
| --- |
| Item | Description |
| 1.0 | Testing Program Details | 7 |
| 1.1 | Equipment Description | 7 |
| 1.1.1 | Equipment Marking Plate | 7 |
| 1.1.2 | Equipment Used During Test | 7 |
| 1.1.3 | Input/Output Ports | 8 |
| 1.1.4 | EUT Internal Operating Frequencies | 8 |
| 1.1.5 | Power Interface | 8 |
| 1.2 | EUT Operation Modes | 9 |
| 1.3 | EUT Configuration Modes | 9 |
| 1.4 | Immunity Performance Criteria | 9 |
| 1.5 | Block Diagram | 11 |
| 1.6 | Compliance Summary | 11 |
| 1.7 | Result Summary | 13 |
| 1.8 | Test Conditions and Results – Conducted Emissions | 14 |
| 1.9 | Test Conditions and Results – Radiated Emissions | 18 |
| 1.10 | Test Conditions and Results – Disturbance Power Emissions | 23 |
| 1.11 | Test Conditions And Results – Limits for Harmonics Current Emissions | 24 |
| 1.12 | Test Conditions And Results – Limits for Voltage Fluctuations and Flicker | 25 |
| 1.13 | Test Conditions and Results – Immunity to electrostatic discharges | 26 |
| 1.14 | Test Conditions and Results – Immunity to RF electromagnetic fields | 27 |
| 1.15 | Test Conditions and Results – Proximity fields from RF wireless communications equipment | 28 |
| 1.16 | Test Conditions and Results – Electrical Fast Transients Immunity | 29 |
| 1.17 | Test Conditions and Results – Surge Immunity | 30 |
| 1.18 | Test Conditions and Results – Conducted Disturbances Immunity | 31 |
| 1.19 | Test Conditions and Results – Power Frequency Magnetic Immunity | 32 |
| 1.20 | Test Conditions and Results – Voltage Dips, Interruptions Immunity | 33 |
| 1.21 | Test Conditions and Results – Electrical transient conduction along supply lines | 34 |

Note: To update all fields in this Test Plan, including the Page numbers on the Index page, click Select All on the Edit menu, and then press F9.

## Equipment Description

### Equipment Marking Plate

(The artwork below may be only a draft. The use of certification marks on a product must be authorized by the respective NCBs that own these marks.)

### Supporting Equipment Used During Test:

| Use* | Product Type | Manufacturer | Model | Comments |
| --- | --- | --- | --- | --- |
| EUT |  |  |  |  |
| Note: * Use one of the following: EUT - Equipment Under Test AE - Auxiliary/Associated Equipment SIM - Simulator (Not Subjected to Test) *Note: Use abbreviations: |

### Input/output Ports:

| Port No. | Name | Type* | Cable Max. >3m | Cable Shielded | Comments (Sip/Sop lines must include description of use) |
| --- | --- | --- | --- | --- | --- |
| 0 | Enclosure | N/E | — | — | None |
| 1 | Mains |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| Supplementary information: *Note: AC = AC Power Port DC = DC Power Port N/E = Non-Electrical Batt=Battery Sip/Sop = Signal Input/output Port PC – Patient-Coupled Cable TP = Telecommunication Ports |

### EUT Internal Operating Frequencies:

| Frequency (MHz) | Description | Frequency (MHz) | Description |
| --- | --- | --- | --- |
| Supplementary information: |

### Power Interface

| Mode No. | Voltage (V) | Current (A) | Power (W) | Frequency (DC/AC-Hz) | Phases (No.) | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| Rated |  |  |  |  |  |  |
| Supplementary information: |

## EUT Operation Modes:

| Mode # | Description |
| --- | --- |
| 1 |  |
| Supplementary information: |

## EUT Configuration Modes

| Mode # | Description |
| --- | --- |
| 1 |  |
| Supplementary information: |

## Basic Safety, Essential Performance and Immunity Pass/Fail Criteria as determined by the Manufacturer

Description of Basic Safety and Essential Performance

Description how the BASIC SAFETY and ESSENTIAL PERFORMANCE were monitored during each test

Immunity Pass/Fail Criteria

## Configuration Block Diagram:

## Compliance Summary

| IEC 60601-1-2 |
| --- |
| Clause | Requirement + Test | Result - Remark | Verdict |

| 4 | GENERAL REQUIREMENTS |
| --- | --- |
| 4.1 | Risk management process for ME equipment and ME systems. |  |  |
| 5 | IDENTIFICATION, MARKING AND DOCUMENTS |
| 5.1 | Equipment specified for use only in shielded location has appropriate marking/warning labels |  |  |
| 5.2 | Accompanying Documents |
| 5.2.1 | Instructions for use |
| 5.2.1.1 | General |
| a) | A statement of the environments the ME equipment will be used. Relevant exclusions, as determined by Risk Analysis, shall also be listed. |  |  |
| b) | The essential performance of ME equipment and a description of what the operator can expect if the Essential Performance is lost or degraded due to EM disturbances. |  |  |
| c) | A warning regarding stacking and location close to other equipment |  |  |
| d) | List of cables, transducers and accessories |  |  |
| e) | A warning that other cables and accessories may negatively affect EMC performance |  |  |
| f) | A statement that portable RF communications equipment. Including antennas, can effect medical electrical equipment. The warning should include a use distance such as “…be used no closer than 30 cm (12 inches) to any part of the [ME EQUIPMENT or ME SYSTEM], including cables specified by the manufacturer” |  |  |
| 5.2.1.2 | Required statement from standard for Class A equipment |  |  |
| 5.2.2 | Technical description |
| 5.2.2.1 | All equipment and systems |
|  | Describe precautions to be taken to prevent adverse events to the Patient and Operator due to electromagnetic disturbances |  |  |
| a) | Compliance information for each test |  |  |
| b) | Statement of any deviations from standards used |  |  |
| c) | Statements to maintain basic safety and essential performance in regards to EMC |  |  |
| 5.2.2.2 | ME Equipment specified for use only in shielded location |
| a) | A warning that equipment should be used only in the specified type of shielded location |  |  |
| b) | Specifications for shielded location including: – minimum RF shielding effectiveness; – for each cable that enters or exits the shielded location, the minimum RF filter attenuation; and – the frequency range(s) over which the specifications apply |  |  |
| c) | Test methods for measurement of RF shielding effectiveness and RF filter attenuation |  |  |
| d) | Required statement(s) from standard and recommended notice for information post at entrance. |  |  |
| 5.2.2.3 | ME Equipment that intentionally receive RF energy shall include: frequency and/or band and bandwidth of receiving section |  |  |
| 5.2.2.4 | ME Equipment that include transmitters shall include frequency and/or band, modulation, and ERP |  |  |
| 5.2.2.5 | Large ME Equipment that are permanently installed |
| a) | A statement that an exemption has been used and that the equipment has not been tested for radiated RF immunity over the entire frequency range 80 MHz to 6 GHz |  |  |
| b) | A warning that the ME equipment has been tested for radiated RF immunity only at selected frequencies |  |  |
| c) | A list of the frequencies and modulations tested |  |  |
| 5.2.2.6 | ME equipment that claim compatibility with HF surgical equipment shall include statement of compatibility and conditions of intended use during HR surgery |  |  |

## Result Summary

| Requirement – Test | Result/Comments | Verdict |
| --- | --- | --- |
| Clause 7 - Emissions |  |
| Classification |  | — |
| Class A or B …: |  | — |
| Group 1 or 2 …: |  | — |
| CISPR 11, 14-1, 32 or ISO 7137 …: |  | — |
| Conducted RF Emissions …: |  |  |
| Radiated RF Emissions …: |  |  |
| Disturbance Power (if applicable) …: |  |  |
| Harmonic Distortion per IEC61000-3-2 (Class A, B, C, D): |  |  |
| Voltage Fluctuations and Flicker per IEC61000-3-3 …: |  |  |
| Clause 8 - Immunity |  |
| Electrostatic Discharges …: |  |  |
| Radiated RF EM Fields and Proximity Wireless fields …: |  |  |
| Electrical Fast Transients and bursts …: |  |  |
| Surges …: |  |  |
| Conducted Disturbances, induced by RF fields …: |  |  |
| Voltage Dips and Interruptions …: |  |  |
| Rated Power-frequency Magnetic Field …: |  |  |

## Test Conditions and Results – Conducted Emissions

| CISPR 11 | TEST: Limits of mains terminal disturbance voltage | Verdict |
| --- | --- | --- |
| Method: The AMN placed 0,8 m from the boundary of the unit under test and bonded to a ground reference plane. This distance was between the closest points of the AMN and the EUT. All other units of the EUT and associated equipment were at least 0,8 m from the AMN. All power was connected to the system through Artificial Mains Network (AMN). Conducted voltage measurements on mains lines were made at the output of the AMN. |  |
| Laboratory Parameters | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 150kHz to 30MHz | Mains |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – Group 1 - Class A |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result* | Average | Result* |
| 0.15 to 0.50 | 79 |  | 66 |  |
| 0.50 to 30 | 73 |  | 60 |  |
| Limits - Group 2 - Class A |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result* | Average | Result* |
| 0.15 to 0.50 | 100 |  | 90 |  |
| 0.50 to 5 | 86 |  | 76 |  |
| 5 to 30 | 90 to 70 |  | 80 to 60 |  |
| Limits - Group 2 - Class A Mains supply currents in excess of 100 A per phase |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result* | Average | Result* |
| 0.15 to 0.50 | 130 |  | 120 |  |
| 0.50 to 5 | 125 |  | 115 |  |
| 5 to 30 | 115 |  | 105 |  |
| Limits - Group 1 and 2 - Class B |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result* | Average | Result* |
| 0.15 to 0.50 | 66 to 56 |  | 56 to 46 |  |
| 0.50 to 5 | 56 |  | 46 |  |
| 5 to 30 | 60 |  | 50 |  |
| Supplementary information: * - The result in tables may be a minimum margin to the limit. EUT powered at one of the Nominal input voltages and frequencies. |

| CISPR 32 | TEST: Limits of mains terminal disturbance voltage | Verdict |
| --- | --- | --- |
| Method: The AMN placed 0,8 m from the boundary of the unit under test and bonded to a ground reference plane. This distance was between the closest points of the AMN and the EUT. All other units of the EUT and associated equipment were at least 0,8 m from the AMN. All power was connected to the system through Artificial Mains Network (AMN). Conducted voltage measurements on mains lines were made at the output of the AMN. |  |
| Laboratory Parameters | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 150kHz to 30MHz | Mains |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – Class A |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result * | Average | Result * |
| 0.15 to 0.50 | 79 |  | 66 |  |
| 0.50 to 30 | 73 |  | 60 |  |
| Limits - Class B |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result * | Average | Result * |
| 0.15 to 0.50 | 66 to 56 |  | 56 to 46 |  |
| 0.50 to 5 | 56 |  | 46 |  |
| 5 to 30 | 60 |  | 50 |  |
| Supplementary information: * - The result in tables may be a minimum margin to the limit. EUT powered at one of the Nominal input voltages and frequencies. |

| CISPR 14 | TEST: Limits of terminal disturbance voltage | Verdict |
| --- | --- | --- |
| Method: The AMN was placed 0,8 m from the boundary of the unit under test and bonded to a ground reference plane. This distance was between the closest points of the AMN and the EUT. All other units of the EUT and associated equipment was at least 0,8 m from the AMN. All power was connected to the system through Artificial Mains Network (AMN). Conducted voltage measurements on mains lines were made at the output of the AMN. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range (see note 1) | Frequency range on each side of line | Measurement Point |
|  | 148.5 kHz to 30MHz | Mains, Load Terminals, Additional Terminals |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – Mains Terminal |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result * | Average | Result * |
| 0.15 to 0.50 | 66 to 56 |  | 59 to 46 |  |
| 0.50 to 5 | 56 |  | 46 |  |
| 5 to 30 | 60 |  | 50 |  |
| Limits - Load /Additional Terminals |
| Frequency (MHz) | Limit dB (µV) |
|  | Quasi-Peak | Result * | Average | Result * |
| 0.15 to 0.50 | 80 |  | 70 |  |
| 0.50 to 5 | 74 |  | 64 |  |
| 5 to 30 | 74 |  | 64 |  |
| Supplementary information: * - The result in this table may be a minimum margin to the limit. EUT powered at one of the Nominal input voltages and frequencies. Note 1: the World Administrative Radio-Communications Conference (WARC) has in 1979 reduced the lower frequency limit in region 1 to 148.5 kHz; for applications falling in the scope of this standard, tests at 150 kHz are considered adequate, since 148,5 kHz falls within the receiver bandwidth. In limit tables 150 kHz used. |

| ISO 7137 | TEST: Limits of mains terminal and I/O disturbance current | Verdict |
| --- | --- | --- |
| Method: Measurements are made with a current probe. Tests are setup per Figure 21-6 of DO-160G. The current probe is installed five centimetres from the EUT. If the EUT connector plus back shell length exceeds five centimetres the probe shall be placed as close to the connector back shell as possible and the position noted. |  |
| Laboratory Parameters | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 150kHz to 152MHz | Mains |
|  | 150kHz to 152MHz | I/O |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – Power Lines |
| Frequency (MHz) | Limit dB (µA) |
|  | Quasi-Peak | Results * |
| 0.15 to 2 | 53 to 20 |  |
| 2 to 30 | 2 |  |
| 30 to 152 | 30 |  |
| Limits – I/O Line |
| Frequency (MHz) | Limit dB (µA) |
|  | Quasi-Peak | Results * |
| 0.15 to 2 | 73 to 40 |  |
| 2 to 30 | 40 |  |
| 30 to 108 | 50 |  |
| 30 to 152 | 40 |  |
| Supplementary information: * - The result in tables may be a minimum margin to the limit. EUT powered at one of the Nominal input voltages and frequencies. |

## Test Conditions and Results – Radiated Emissions

| CISPR 11 | TEST: Limits for radiated disturbance 30 MHz –1 GHz | Verdict |
| --- | --- | --- |
| Method: Measurements were made in a 10-meter semi-anechoic chamber or Open Area Test Site that complies to CISPR 16. Preliminary (peak) measurements were performed at an antenna to EUT separation distance of listed below. The EUT was rotated 360° about its azimuth with the receive antenna located at various heights in horizontal and vertical polarities. Final measurements (quasi-peak) were then performed by rotating the EUT 360° and adjusting the receive antenna height from 1 to 4 m. All frequencies were investigated in both horizontal and vertical antenna polarity, where applicable. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 30MHz – 1GHz | 10 m measurement distance*** |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – Group 1 Class A |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Quasi-Peak | Results * |
| 30 to 230 | 40 |  |
| 230 to 1000 | 47 |  |
| Limits – Group 1 Class B |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Quasi-Peak | Results * |
| 30 to 230 | 30 |  |
| 230 to 1000 | 37 |  |
| Limits – Group 2 Class A |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Quasi-Peak | Results * |
| 30 to 1000 | See standard |  |
| Limits – Group 2 Class B |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Quasi-Peak | Average** | Results * |
| 30 to 80,872 | 30 | 25 |  |
| 80,872 to 81,848 | 50 | 45 |  |
| 81,848 to 134,786 | 30 | 25 |  |
| 134,786 to 136,414 | 50 | 45 |  |
| 136,414 to 230 | 30 | 25 |  |
| 230 to 1000 | 37 | 32 |  |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. * - The result in this table may be a minimum margin to the limit. ** - The Average Limits Apply To Magnetron Driven Equipment Only. *** - Test distance used shall follow guidance from CISPR 11 |

| CISPR 32 | TEST: Limits for radiated disturbance 30 MHz –XX GHz** | Verdict |
| --- | --- | --- |
| Method: Measurements were made in a 10-meter semi-anechoic chamber or Open Area Test Site that complies to CISPR 16. Preliminary (peak) measurements were performed at an antenna to EUT separation distance of 10 meter below 1GHz and 3 meter above 1GHz. The EUT was rotated 360° about its azimuth with the receive antenna located at various heights in horizontal and vertical polarities. Final measurements (quasi-peak detector below 1GHz and average detector above 1GHz) were then performed by rotating the EUT 360° and adjusting the receive antenna height from 1 to 4 m. All frequencies were investigated in both horizontal and vertical antenna polarity, where applicable. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 30MHz – 1GHz | 10 m measurement distance |
|  | 1GHz – 6GHz | 3 m measurement distance |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – Class A |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Detector | Results * |
| 30 to 230 | 40 Quasi-Peak |  |
| 230 to 1000 | 47 Quasi-Peak |  |
| 1000 to 3000 | 56 Average, 76 Peak |  |
| 3000 to 6000 | 76 Average, 80 Peak |  |
| Limits – Class B |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Detector | Results * |
| 30 to 230 | 30 Quasi-Peak |  |
| 230 to 1000 | 37 Quasi-Peak |  |
| 1000 to 3000 | 50 Average, 70 Peak |  |
| 3000 to 6000 | 54 Average, 74 Peak |  |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. * - The result in this table may be a minimum margin to the limit. ** - Input highest test frequency based on internal frequency used in the ME equipment |

| CISPR 14-1 | TEST: Limits for radiated disturbance 30 MHz –1 GHz | Verdict |
| --- | --- | --- |
| Method: Measurements were made in a 10-meter semi-anechoic chamber or Open Area Test Site that complies to CISPR 16. Preliminary (peak) measurements were performed at an antenna to EUT separation distance of 10 meter. The EUT was rotated 360° about its azimuth with the receive antenna located at various heights in horizontal and vertical polarities. Final measurements (quasi-peak) were then performed by rotating the EUT 360° and adjusting the receive antenna height from 1 to 4 m. All frequencies were investigated in both horizontal and vertical antenna polarity, where applicable. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 30MHz – 1GHz | 10 m measurement distance |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits – CISPR 14-1 |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Quasi-Peak | Results * |
| 30 to 230 | 30 |  |
| 230 to 1000 | 37 |  |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. * -The result in this table may be a minimum margin to the limit. |

| ISO 7137 | TEST: Limits for radiated disturbance 100 MHz –6 GHz | Verdict |
| --- | --- | --- |
| Method: Measurements were made in a shielded enclosure. RF absorber shall be placed above, behind, and on both sides of test setup boundary, from ceiling to ground plane. The absorber shall extend ≥50 cm in front of the ground plane. RF absorber shall be placed behind the test antenna, from ceiling to floor. The distance between the absorber and the antenna shall be ≥30 cm. Linearly polarized antennas are required for radiated tests. Measure radiated emissions using both vertically and horizontally polarized orientations. Other requirements in DO-160G, clause 21.5, shall be followed. If radiated RF emissions are conducted in a reverberation chamber, the procedures in DO-160G, clause 21.6, shall be followed. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range on each side of line | Measurement Point |
|  | 100MHz – 1GHz | 0.9 m from the front of the test table |
|  | 1GHz – 6GHz |  |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits |
| Frequency (MHz) | Limit dB (µV/m) |
|  | Detector | Results * |
| 100 to 108 | 44.6 to 45.2 (QP) |  |
| 108 to 152 | 35 to 37.5 (QP) |  |
| 152 to 960 | 47.5 to 60.2 (QP) |  |
| 960 to 1215 | 50.3 to 52 (AV) |  |
| 1215 to 1525 | 61.9 to 63.5 (AV) |  |
| 1525 to 1680 | 53.5 to 54 (AV) |  |
| 1680 to 6000 | 64 to 73 (AV) |  |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. * -The result in this table may be a minimum margin to the limit. |

## Test Conditions and Results – Disturbance Power Emissions

| CISPR 14-1 | TEST: Limits of disturbance power | Verdict |
| --- | --- | --- |
| Method: Measurements were made on a ground plane that extends 1-meter minimum beyond all sides of the system under test. All power was connected to the system through Line Impedance Stabilization Networks (LISN). The lead to be measured on is stretched in a straight line for a distance sufficient to accommodate the absorbing clamp, and to permit the necessary measuring adjustment of position for tuning. The clamp is placed around the lead so as to measure a quantity proportional to the disturbance on the lead. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 15 to 35 °C | °C |
| Relative Humidity | 30 to 60 % | % |
| Fully configured sample scanned over the following frequency range | Frequency range | Measurement point |
|  | 30 MHz to 300 MHz | Mains |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Limits disturbance power on terminals |
| Frequency 30 – 300 MHz | dB(pW) |
|  | Quasi-Peak | Result * | Average | Result * |
| Household and Tools < 700 W | 45 to 55 |  | 35 to 45 |  |
| 700 W < Tools < 1000 W | 49 to 59 |  | 39 to 49 |  |
| Tools >1000 W | 55 to 65 |  | 45 to 55 |  |
| Supplementary information: * - The result in this table may be a minimum margin to the limit. EUT powered at one of the Nominal input voltages and frequencies. |

## Test Conditions and Results – Harmonic Distortion

| 61000-3-2 | TEST: Limits for Harmonic current emissions (IEC 61000-3-2:2005 A1:2008 A2:2009) | Verdict |
| --- | --- | --- |
| Method: This test consists on the measurement of harmonics components of the input current which may be produced by equipment having an input current up to and including 16 A per phase, and intended to be connected to public low-voltage distribution systems. The equipment is tested under specified conditions of operation. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 15 to 35 °C | °C |
| Relative Humidity | 30 to 60 % | % |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Classification of Equipment : | Class |
| Supplementary information: |

## Test Conditions and Results – Voltage Fluctuations and Flicker

| 61000-3-3 | TEST: Limitation of Voltage Fluctuations And Flicker (IEC 61000-3-3:2013) | Verdict |
| --- | --- | --- |
| Method: The test circuit consists of a test supply voltage, reference impedance, the equipment under test and a flicker meter compliant with IEC 60868. The equipment shall be tested in the condition in which the manufacturer supplies it. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 15 to 35 °C | °C |
| Relative Humidity | 30 to 60 % | % |
| Equipment mode | Power interface mode : |  |
|  | EUT configurations mode : |  |
|  | Operation mode : |  |
| Control Method of Equipment (see below) : |  |
| 1 - without additional conditions |
| 2 - switched manually, or switched automatically more frequently than twice per day, and also has either a delayed restart (the delay being not less than a few tens of seconds), or manual restart, after a power supply interruption. |
| 3 - attended while in use (for example: hair dryers, vacuum cleaners, kitchen equipment such as mixers, garden equipment such as lawn mowers, portable tools such as electric drills), or switched on automatically, or is intended to be switched on manually, no more than twice per day, and also has either a delayed restart (the delay being not less than a few tens of seconds) or manual restart, after a power supply interruption. |
| Supplementary information: |

## Test Conditions and Results – Immunity to Electrostatic Discharges

| 61000-4-2 | TEST: Electrostatic discharges (IEC 61000-4-2:2008) | Verdict |
| --- | --- | --- |
| Method: The test is intended to demonstrate the immunity of equipment subjected to static electricity discharges from operators directly and to adjacent objects. The table top equipment under test is placed on a wooden table, 0.8 m high, standing on the ground reference plane. A horizontal coupling plane (HCP), 1.6 x 0.8 m, is placed on the table. The EUT and the cables are isolated from the coupling plane by an insulating support 0.5 mm thick. The floor standing equipment is isolated from the ground reference plane by an insulating support about 0.1 m thick. The vertical coupling plane (VCP) of dimensions 0.5 m x 0.5 m is placed parallel to, and positioned at a distance of 0.1 m from, the EUT. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 15 to 35 °C | °C |
| Relative Humidity | 30 to 60 % | % |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Test Levels |
| Discharge type | Discharge Level (kV) | Number of discharges per location (each polarity) |
|  | Positive | Negative |  |
| Air – Direct | 2,4,8,15 | 2,4,8,15 | 10 |
| Contact – Direct | 8 | 8 | 10 |
| Contact – Indirect | 8 | 8 | 10 |
| Discharge location | See photo documentation of the test set-up All external locations accessible by hand, Horizontal plate (HCP) Vertical coupling plate (VCP) |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies |

## Test Conditions and Results - Immunity to Radio Frequency Electromagnetic Fields

| 61000-4-3 | TEST: RF electromagnetic fields (IEC 61000-4-3:2006 A1:2007 A2:2010) | Verdict |
| --- | --- | --- |
| Method: The test allows estimating of the radiated immunity of electrical and electronic equipment to electromagnetic disturbances coming from intended radio-frequency (RF) transmitters in the frequency range 80 MHz to 2500 MHz. The interference is applied on the enclosure of the equipment by using transmitting antennas. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 15 to 35 °C | °C |
| Relative Humidity | 30 to 60 % | % |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Test specifications |
| Calibration Requirements | Uniform field area (UFA) | 1.5 m x 1.5 m, 16 points with a minimum UFA size 0.5 m x 0.5 m |
|  |  | 75 % of calibration points within specifications if UFA is larger than 0.5 m x 0.5 m. 100 % (all 4 points) in the specifications for 0.5 x 0.5 m UFA |
| Frequency bandwidth | 80 MHz to 2700 MHz |
| Modulation | 80% AM at 1kHz* |
| Level | Professional Healthcare Environment | 3 V/m |
|  | Home Healthcare Environment | 10 V/m |
| Frequency step | 1% |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. Dwell time minimum 1 s. Actual dwell time noted in results table. Actual test level noted in results table. Note * - Testing may be performed at other modulation frequencies identified by the Risk Management Process. |

## Test Conditions and Results – Proximity fields from RF wireless communications equipment

| 61000-4-3 | TEST: Proximity fields from RF wireless communications equipment (IEC 61000-4-3:2006 A1:2007 A2:2010) | Verdict |
| --- | --- | --- |
| Method: The test allows estimating of the radiated immunity of electrical and electronic equipment to electromagnetic disturbances coming from RF communications equipment. The interference is applied on the enclosure of the equipment by using transmitting antennas. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 15 to 35 °C | °C |
| Relative Humidity | 30 to 60 % | % |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Test specifications |
| Calibration Requirements | Uniform field area (UFA) | 1.5 m x 1.5 m, 16 points with a minimum UFA size / 0.5 m x 0.5 m |
|  |  | 75 % of calibration points within specifications if UFA is larger than 0.5 m x 0.5 m. 100 % (all 4 points) in the specifications for 0.5 x 0.5 m UFA |

| Frequency Range and Level: RF wireless communication equipment |
| --- |
| Test Frequency (MHz) | Modulation | Immunity Level (V/m) |
| 385 | **Pulse Modulation: 18Hz | 27 |
| 450 | *FM + 5Hz deviation: 1kHz sine | 28 |
| 710 745 780 | **Pulse Modulation: 217Hz | 9 |
| 810 870 930 | **Pulse Modulation: 18Hz | 28 |
| 1720 1845 1970 | **Pulse Modulation: 217Hz | 28 |
| 2450 | **Pulse Modulation: 217Hz | 28 |
| 5240 5500 5785 | **Pulse Modulation: 217Hz | 9 |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. Dwell time minimum 1 s. Actual dwell time noted in results table. Note * - As an alternative to FM modulation, 50 % pulse modulation at 18 Hz may be used because while it does not represent actual modulation, it would be worst case. Note** - The carrier shall be modulated using a 50 % duty cycle square wave signal. |

## Test Conditions and Results – Electrical Fast Transients

| 61000-4-4 | TEST: Fast Transients – (IEC61000-4-4:2012) | Verdict |
| --- | --- | --- |
| Method: Measurements were made on a ground plane that extends 1-meter minimum beyond all sides of the system under test. Mains power tests were conducted with the product connected to a Coupling/Decoupling Network (CDN). I/O lines were tested in a Capacitive Coupling Clamp. One of each unique interface was tested for a period of one (1) minute per polarity. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample subject to the levels shown below. | Measurement Point |
|  | Input a.c. Power Ports Input d.c. Power Ports Signal Input/Output Ports |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Applied Level |
| Application Point | (kV) | Coupling Method | Repetition Frequency (kHz) |
| Input a.c. Power Ports | ±2 | Direct Injection | 100 |
| Input d.c. Power Ports | ±2 | Direct Injection | 100 |
| Signal Input/Output Ports* | ±1 | Capacitive Clamp | 100 |
| Input d.c. Power ports (transportation) | See ISO 7637-2 |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. Note* - SIP/SOPS whose maximum cable length is less than 3m are excluded |

## Test Conditions and Results – Surge Immunity

| 61000-4-5 | TEST: Surge Immunity Test – (IEC61000-4-5:2005) | Verdict |
| --- | --- | --- |
| Method: Mains power tests were conducted with the product connected to a Coupling/ Decoupling Network (CDN). The test voltage was increased from the lowest indicated level up to the maximum level. Five (5) positive surges and five (5) negative surges were applied at each of phases of the a.c. waveform: 0°, 90°, 180° and 270°. Each surge was applied 60 seconds after the previous surge. Signal and Telecommunications ports were subject to five (5) positive and five (negative) surges applied through the appropriate Coupling/Decoupling Network (CDN). |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample subject to the levels shown below. | Measurement Point |
|  | Input AC and DC Power Ports |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Applied Level |
| Application Point | [kV] | Required Surge Waveform |
| Input Power Ports | 0.5 and 1.0 (Line to Line) | Combination Wave (2µs x 50µs Voltage, 8µs x 20µs Current) |
|  | 0.5, 1.0 and 2.0 (Line to Earth) | Combination Wave (2µs x 50µs Voltage, 8µs x 20µs Current) |
| Signal input/output* | 2.0 (Line to Earth) | Combination Wave (2µs x 50µs Voltage, 8µs x 20µs Current) |
| ME EQUIPMENT and ME SYSTEMS that do not have a surge protection device in the primary power circuit may be tested only at 2 kV line(s) to earth and 1 kV line(s) to line(s). |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. Note* - Applicable to output lines intended to connect directly to outdoor cables |

## Test Conditions and Results – Conducted Disturbances Immunity

| 61000-4-6 | TEST: RF Continuous Conducted – (IEC61000-4-6:2013) | Verdict |
| --- | --- | --- |
| Method: Measurements were made on a ground plane that extends 0.5-meter minimum beyond all sides of the system under test. The EUT was located 10cm above the reference ground plane and any associated I/O cables attached to the EUT were located between 30mm and 50mm above the ground plane. The indicated field was pre-calibrated prior to placement of the system under test. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Test Specifications: | Frequency range | Measurement Point |
| Fully configured sample scanned over the following frequency range | 150kHz to 80MHz | Input a.c. Power Ports Input d.c. Power Ports Signal Input/output Ports Patient Connected Ports |
| Level | Home Healthcare Environment | 3 V RMS outside the ISM band, 6 V RMS in the ISM and amateur radio bands |
|  | Professional Healthcare Environment | 3 V RMS outside the ISM band, 6 V RMS in the ISM band |
| Frequency step | 1% |
| Modulation | 80% Am at 1kHz* |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequencies. Dwell time minimum 1 s. Note * - Testing may be performed at other modulation frequencies identified by the Risk Management Process. |

## Test Conditions and Results – Power- Frequency Magnetic Fields

| 61000-4-8 | TEST: Power-frequency magnetic field – (IEC61000-4-8:2009) | Verdict |
| --- | --- | --- |
| Method: Measurements were made on a ground plane that extends 1-meter minimum beyond sides of the system under test. Table top EUT is located 80cm above the reference ground plane and floor-standing EUT is located 10cm above the reference ground plane. The indicated field was pre-calibrated prior to placement of the EUT under test. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Fully configured sample tested at the power line frequency (See Note 1) | Frequency | Application Point |
|  | 50Hz and 60 Hz | Enclosure |
| Test Level | 30 (A/m) |
| Supplementary information: EUT powered at one of the Nominal input voltages and frequency of either 50Hz or 60Hz. During the test, the frequency of the field and line frequency of the ME equipment shall be the same. |

## Test Conditions and Results – Voltage Dips, Interruptions, and Variations

| 61000-4-11 | TEST: Voltage Dips and Interruptions – (IEC61000-4-11:2004) | Verdict |
| --- | --- | --- |
| Method: The product was subjected to voltage dips and interruptions. Testing was performed with the product connected directly to a generator capable of simulating the voltage drops and interrupts as described. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured subjected to the levels indicated below. | Measurement Point |
|  | Input A.C. Power Ports |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Applied Levels |
| Voltage Dips % UT | Cycles | Sync Angle (degrees) |
| >95 | 0.5* | 0,45,90,135,180,225,270,315 |
| >95 | 1 | 0 |
| 30 | 25 (50Hz) 30 (60Hz) | 0 |
| Voltage Interruption % UT | Cycles | Sync Angle [degrees] |
| >95 | 250 (50Hz),300 (60Hz) | 0 |
| Supplementary information: If the Rated voltage range <25% of the lowest rated input voltage, one rated input voltage. Otherwise, minimum and maximum rated voltage. EUT powered at one of the Nominal input frequencies. ME EQUIPMENT and ME SYSTEMS with power input voltage selection by transformer taps shall be tested at only one tap setting. Note* - Only applicable to ME equipment with single phase a.c. mains |

## Test Conditions and Results – Electrical transient conduction along supply lines

| ISO 7637-2 | TEST: Electrical transient conduction along supply lines | Verdict |
| --- | --- | --- |
| Method: The DUT shall be placed on a non-conductive low relative permittivity ( εr ≤ 1.4) support with a thickness of (50 ± 5) mm. Grounding of the DUT case to the ground plane shall reflect the vehicle installation. Preferably, the load simulator shall be placed directly on the ground plane. If the load simulator has a metallic case, this case shall be bonded to the ground plane. The peak voltage Us shall be adjusted to be the required test level with the tolerances of +10 % and 0 %. |  |
| Laboratory Parameters: | Required prior to the test | During the test |
| Ambient Temperature | 10 to 40 °C | °C |
| Relative Humidity | 10 to 90 % | % |
| Fully configured sample subject to the levels shown below. |
| Equipment mode | Power interface mode |  |
|  | EUT configurations mode |  |
|  | Operation mode |  |
| Applied Level |
| Application Point | Test pulse and level | Coupling Method |
| Input d.c. Power ports (transportation) | 1 (-150V) | Direct Injection |
|  | 2a (+112V) | Direct Injection |
|  | 2b (10V for 12V system; 20V for 24V system) | Direct Injection |
|  | 3a (-220V for 12V system; -300V for 24Vsystem) | Direct Injection |
|  | 3b (+150V for 12V system; +300 for 24V system) | Direct Injection |
| Supplementary information: For ME EQUIPMENT and ME SYSTEMS intended to be installed in passenger cars and light commercial vehicles including ambulances fitted with 12 V electrical systems or commercial vehicles including ambulances fitted with 24 V electrical systems. |

End of the Test Plan